## FUS-Tx-BrainBBB744-HEAD
Protocol Name: | FUS-Tx-BrainBBB744-HEAD
| Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement
Field Name \ Series Name | ****CALIBRATION**** | Tracking | Tracking Prescan | ManualCalibration | CentralFrequency | MovementDetection | ****PLANNING**** | T2 FSE REFORMAT | T1w SPGR REFORMAT | ****SONICATION**** | T2STAR-SE | T2STAR-ME | TMAP | MEMP | Tx-BBB744-3T-750-DV26.0-001
Series Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Imaging Mode | 2D | 2D | 2D | 2D | 2D | 2D | 2D | 2D | 3D | 2D | 2D | 2D | 2D | 2D | 2D
Pulse Sequence | Gradient Echo | Gradient Echo | Gradient Echo | FSE-XL | Gradient Echo | Gradient Echo | Gradient Echo | FSE-XL | SPGR | Gradient Echo | SPGR | Multi-Echo FSPGR | SPGR | Multi-Echo FSPGR | Gradient Echo
PSD Name | track | track | NA | NA | NA | NA | track | NA | NA | track | NA | fgre_memp | NA | fgre_memp | track
Imaging Options | Seq, Fast | Seq, Fast | Seq, Fast | EDR, TRF, Fast, FR | Fast | NPW, Seq, Fast | Seq, Fast | EDR, TRF, Fast, ZIP512, FR | EDR, Fast | Seq, Fast | Seq, Fast, MPh | FC, Seq, Fast, MPh | Seq, Fast, MPh | Seq, Fast, MPh | Seq, Fast
FOV | 480 | 48 | 24 | 40 | 34 | 24 | 480 | 27 | 29 | 480 | 26 | 27 | 28 | 28 | 480
Phase FOV | 1 | 1 | 1 | 1 | 1 | NA | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Slices/Slabs | 1 | 1 | 3 | 1 | 1 | 3 | 1 | 70 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Scan Locs | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slices per Slab | NA | NA | NA | NA | NA | NA | NA | NA | 128 | NA | NA | NA | NA | NA | NA
Slice Thickness | 5 | 5 | 5 | 3 | 4 | 5 | 5 | 2 | 1.5 | 5 | 3 | 3 | 3 | 3 | 5
NPW_FACTOR | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
No Slab Wrap | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slice Resolution | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Spacing | 0 | 0 | 5 | 1 | 1 | 5 | 0 | 0 | NA | 0 | 1 | 1.5 | 1.5 | 1.5 | 0
Plane | 3-PLANE | 3-PLANE | 3-PLANE | SAGITTAL | AXIAL | 3-PLANE | 3-PLANE | AXIAL | SAGITTAL | 3-PLANE | AXIAL | AXIAL | AXIAL | AXIAL | 3-PLANE
Freq DIR | Unswap | Unswap | Unswap | S/I | R/L | Unswap | Unswap | A/P | S/I | Unswap | R/L | R/L | A/P | A/P | Unswap
TR | 20 | 20 | NA | 2500 | 100 | NA | 20 | 6700 | NA | 20 | NA | 55 | NA | 28 | 20
Auto TR | Off | Off | Off | Auto | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off
TE | Minimum | Minimum | NA | 85 | 5 | NA | Minimum | 95.4 | Min Full | Minimum | Min Full | Min Full | Min Full | Min Full | Minimum
TE2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Number of Echoes | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 6 | 1 | 5 | 1
Number of Shots | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Echo Train Length | NA | NA | NA | 16 | NA | NA | NA | 16 | NA | NA | NA | NA | NA | NA | NA
Flip Angle | NA | NA | NA | 111 | 30 | NA | NA | 111 | 14 | NA | 15 | 20 | 15 | 15 | NA
Auto Flip Angle | NA | NA | NA | 1 | NA | NA | NA | 1 | NA | NA | NA | NA | NA | NA | NA
Mat Freq | 1024 | 1024 | 256 | 256 | 256 | 256 | 1024 | 384 | 288 | 1024 | 256 | 256 | 256 | 256 | 1024
Mat Phase | 2 | 2 | 128 | 192 | 128 | 128 | 2 | 352 | 288 | 2 | 256 | 256 | 128 | 128 | 2
NEX | NA | NA | 1 | 1 | 1 | 2 | NA | 1 | 1 | NA | 1 | 1 | 1 | 1 | NA
BW | 62.5 | 62.5 | NA | 31.25 | 31.25 | NA | 62.5 | 31.25 | 25 | 62.5 | 3.97 | 19.23 | 5.68 | 35.71 | 62.5
BW2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
TI | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
INVERTF | NA | NA | NA | Normal (S) | NA | NA | NA | Normal (A) | NA | NA | NA | NA | NA | NA | NA
CV | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=480.00)(USERCV8=480.00)(USERCV9=480.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=16.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV6=0.00)(USERCV7=0.00)(USERCV8=0.00)(USERCV9=0.00)(USERCV15=0.00)(USERCV18=0.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV6=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV6=0.00)(USERCV7=0.00)(USERCV8=0.00)(USERCV9=0.00)(USERCV15=0.00)(USERCV18=0.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV38=0.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=0.00)(USERCV23=100.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV18=1.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV18=1.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00)
Acceleration Phase | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acceleration Slice | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Delay | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Minimum | Minimum | Minimum | Minimum | NA
Phases Per Location | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 5 | 5 | 6 | 6 | NA
Acquisition Order | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0 | 0 | 0 | 0 | NA
PURE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anatomical Region | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain
Coil | 3P 2ch Head | 3P 8ch Tracking | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head
CLINICALCOIL | C-FUS 2ch Head | C-FUS 8ch Tracking | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head
COILCOMPONENT | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec
CoilReg | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CoilPlug | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Shim | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto
EXCITATIONMODE | Non-Selective | Non-Selective | Selective | Selective | Selective | Selective | Non-Selective | Selective | Selective | Non-Selective | Selective | Selective | Selective | Selective | Non-Selective
SL3PLANE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1
SL3PLANE2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE3 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SLICEORDER | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SPCPERPLANE1 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE2 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE3 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PHASECORR | No | No | No | Yes | No | No | No | Yes | No | No | No | No | No | No | No
SAVEORIGINAL | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## FUS-Tx-BrainBBB744-BODY
Protocol Name: | FUS-Tx-BrainBBB744-BODY
| Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement
Field Name \ Series Name | ****CALIBRATION**** | Tracking | Tracking Prescan | ManualCalibration | CentralFrequency | MovementDetection | ****PLANNING**** | T2 FRFSE REFORMAT | T1w SPGR REFORMAT | ****SONICATION**** | T2STAR-SE | T2STAR-ME | TMAP | MEMP | Tx-BBB744-3T-750-DV26.0-001
Series Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
Imaging Mode | 2D | 2D | 2D | 2D | 2D | 2D | 2D | 2D | 3D | 2D | 2D | 2D | 2D | 2D | 2D
Pulse Sequence | Gradient Echo | Gradient Echo | Gradient Echo | FSE-XL | Gradient Echo | Gradient Echo | Gradient Echo | FSE-XL | SPGR | Gradient Echo | SPGR | Multi-Echo FSPGR | SPGR | Multi-Echo FSPGR | Gradient Echo
PSD Name | track | track | NA | NA | NA | NA | track | NA | NA | track | NA | fgre_memp | NA | fgre_memp | track
Imaging Options | Seq, Fast | Seq, Fast | Seq, Fast | EDR, TRF, Fast, FR | Fast | NPW, Seq, Fast | Seq, Fast | NPW, EDR, TRF, Fast, ZIP512, FR | EDR, Fast | Seq, Fast | Seq, Fast, MPh | FC, Seq, Fast, MPh | Seq, Fast, MPh | Seq, Fast, MPh | Seq, Fast
FOV | 480 | 48 | 24 | 40 | 34 | 24 | 480 | 27 | 29 | 480 | 27 | 27 | 28 | 28 | 480
Phase FOV | 1 | 1 | 1 | 1 | 1 | NA | 1 | NA | 1 | 1 | 1 | 1 | 1 | 1 | 1
Slices/Slabs | 1 | 1 | 3 | 1 | 1 | 3 | 1 | 70 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Scan Locs | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slices per Slab | NA | NA | NA | NA | NA | NA | NA | NA | 128 | NA | NA | NA | NA | NA | NA
Slice Thickness | 5 | 5 | 5 | 3 | 4 | 5 | 5 | 2 | 1.5 | 5 | 3 | 3 | 3 | 3 | 5
NPW_FACTOR | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
No Slab Wrap | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slice Resolution | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Spacing | 0 | 0 | 5 | 1 | 1 | 5 | 0 | Intleave | NA | 0 | 1 | 1.5 | 1.5 | 1.5 | 0
Plane | 3-PLANE | 3-PLANE | 3-PLANE | SAGITTAL | AXIAL | 3-PLANE | 3-PLANE | AXIAL | SAGITTAL | 3-PLANE | AXIAL | AXIAL | AXIAL | AXIAL | 3-PLANE
Freq DIR | Unswap | Unswap | Unswap | S/I | R/L | Unswap | Unswap | A/P | S/I | Unswap | R/L | R/L | A/P | A/P | Unswap
TR | 20 | 20 | NA | 2500 | 100 | NA | 20 | 6700 | NA | 20 | NA | 55 | NA | 28 | 20
Auto TR | Off | Off | Off | Auto | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off
TE | Minimum | Minimum | NA | 85 | 5 | NA | Minimum | 98.7 | Min Full | Minimum | Min Full | Min Full | Min Full | Min Full | Minimum
TE2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Number of Echoes | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 6 | 1 | 5 | 1
Number of Shots | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Echo Train Length | NA | NA | NA | 16 | NA | NA | NA | 16 | NA | NA | NA | NA | NA | NA | NA
Flip Angle | NA | NA | NA | 111 | 30 | NA | NA | 111 | 14 | NA | 15 | 20 | 15 | 15 | NA
Auto Flip Angle | NA | NA | NA | 1 | NA | NA | NA | 1 | NA | NA | NA | NA | NA | NA | NA
Mat Freq | 1024 | 1024 | 256 | 256 | 256 | 256 | 1024 | 384 | 288 | 1024 | 256 | 256 | 256 | 256 | 1024
Mat Phase | 2 | 2 | 128 | 192 | 128 | 128 | 2 | 352 | 288 | 2 | 256 | 256 | 128 | 128 | 2
NEX | NA | NA | 1 | 1 | 1 | 2 | NA | 1 | 1 | NA | 1 | 1 | 1 | 1 | NA
BW | 62.5 | 62.5 | NA | 31.25 | 31.25 | NA | 62.5 | 31.25 | 25 | 62.5 | 3.97 | 19.23 | 5.68 | 35.71 | 62.5
BW2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
TI | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
INVERTF | NA | NA | NA | Normal (S) | NA | NA | NA | Normal (A) | NA | NA | NA | NA | NA | NA | NA
CV | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=480.00)(USERCV8=480.00)(USERCV9=480.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=16.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV6=0.00)(USERCV7=0.00)(USERCV8=0.00)(USERCV9=0.00)(USERCV15=0.00)(USERCV18=0.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV6=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV6=0.00)(USERCV7=0.00)(USERCV8=0.00)(USERCV9=0.00)(USERCV15=0.00)(USERCV17=1.00)(USERCV18=0.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV38=0.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=0.00)(USERCV23=100.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV18=1.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV18=1.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00)
Acceleration Phase | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acceleration Slice | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Delay | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Minimum | Minimum | Minimum | Minimum | NA
Phases Per Location | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 5 | 5 | 6 | 6 | NA
Acquisition Order | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0 | 0 | 0 | 0 | NA
PURE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anatomical Region | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain
Coil | BODY | FUS 8ch Tracking | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY
CLINICALCOIL | C-GE_BODY | C-GE_FUS 8ch Tracking | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY
COILCOMPONENT | GE T/R body coil | FUS 8ch Tracking Array by InSightec | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil
CoilReg | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CoilPlug | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Shim | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto
EXCITATIONMODE | Non-Selective | Non-Selective | Selective | Selective | Selective | Selective | Non-Selective | Selective | Selective | Non-Selective | Selective | Selective | Selective | Selective | Non-Selective
SL3PLANE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1
SL3PLANE2 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE3 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SLICEORDER | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SPCPERPLANE1 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | Intleave | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE2 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | Intleave | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE3 | 0 | 0 | 5 | 0 | 0 | 5 | 0 | Intleave | 0 | 0 | 0 | 0 | 0 | 0 | 0
PHASECORR | No | No | No | Yes | No | No | No | Yes | No | No | No | No | No | No | No
SAVEORIGINAL | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## FUS-PrePost-BrainBBB744
Protocol Name: | FUS-PrePost-BrainBBB744
| Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement
Field Name \ Series Name | Localizer T1 | 3D CUBE T2 | qT1_FA12_ROT | qT1_FA2_ROT | qT1_FA5_ROT | qT1_FA8_ROT | qT1_FA16_ROT | B1MAP_ROTcenter | GRE-ME_T2STAR_ROT | GRE-ME_T2STAR_FOV23_ROT | T1w_FSPGR_Brain AX | B1MAP_BrainCenter | Place Holder - Injection, Perfusion | Place Holder - DWI ~2:30min | T2 FLAIR CUBE | T1 CUBE | T1w_FSPGR_Brain AX + Gd | T1w_FSPGR_Brain AX + Gd_15min
Series Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Imaging Mode | 2D | 3D | 3D | 3D | 3D | 3D | 3D | 2D | 2D | 2D | 3D | 2D | 2D | 2D | 3D | 3D | 3D | 3D
Pulse Sequence | Gradient Echo | Cube | SPGR | SPGR | SPGR | SPGR | SPGR | Gradient Echo | Multi-Echo FSPGR | Multi-Echo FSPGR | SPGR | Gradient Echo | Gradient Echo | Spin Echo | Cube T2 FLAIR | Cube | SPGR | SPGR
PSD Name | NA | NA | NA | NA | NA | NA | NA | B1MAP | NA | NA | NA | B1MAP | NA | NA | NA | NA | NA | NA
Imaging Options | Seq, Fast | FC, EDR, Fast, ZIP512, ZIP2, FR, ARC | EDR, Fast, Asset | EDR, Fast, Asset | EDR, Fast, Asset | EDR, Fast, Asset | EDR, Fast, Asset | EDR, Fast | EDR, Fast, Asset | EDR, Fast, Asset | Fast, Asset | EDR, Fast | MPh, EPI, Asset | EDR, EPI, DIFF | EDR, Fast, ZIP512, ZIP2, ARC, IrP | EDR, Fast, ZIP512, ZIP2, FR, ARC | Fast, Asset | Fast, Asset
FOV | 44 | 24 | 22 | 22 | 22 | 22 | 22 | 25 | 25 | 23 | 25.6 | 25 | 22 | 25 | 25.6 | 25.6 | 25.6 | 25.6
Phase FOV | 1 | 1 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | NA | 0.72 | 0.8 | 0.75 | NA | 1 | 1 | 1 | 1 | 0.75 | 0.75
Slices/Slabs | 33 | 1 | 1 | 1 | 1 | 1 | 1 | 15 | 37 | 40 | 1 | 15 | 20 | 200 | 1 | 1 | 1 | 1
Scan Locs | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slices per Slab | NA | 110 | 40 | 40 | 40 | 40 | 40 | NA | NA | NA | 110 | NA | NA | NA | 136 | 140 | 110 | 110
Slice Thickness | 7 | 1.4 | 2 | 2 | 2 | 2 | 2 | 5 | 2.5 | 2.5 | 1.5 | 5 | 5 | 2.9 | 1.4 | 1.2 | 1.5 | 1.5
NPW_FACTOR | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
No Slab Wrap | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slice Resolution | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Spacing | 3 | NA | NA | NA | NA | NA | NA | 5 | 0 | 0 | NA | 5 | 0 | 0 | NA | NA | NA | NA
Plane | 3-PLANE | SAGITTAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | AXIAL | SAGITTAL | SAGITTAL | AXIAL | AXIAL
Freq DIR | Unswap | S/I | A/P | A/P | A/P | A/P | A/P | A/P | A/P | A/P | A/P | A/P | R/L | A/P | S/I | S/I | A/P | A/P
TR | NA | 2500 | NA | NA | NA | NA | NA | NA | 900 | 900 | NA | NA | 1500 | 10000 | 4000 | 600 | NA | NA
Auto TR | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off
TE | NA | 92.9 | Min Full | Min Full | Min Full | Min Full | Min Full | NA | Min Full | Min Full | Min Full | NA | 30 | NA | 105 | Minimum | Min Full | Min Full
TE2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Number of Echoes | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 8 | 8 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Number of Shots | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1 | 1 | NA | NA | NA | NA
Echo Train Length | NA | 140 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 200 | 18 | NA | NA
Flip Angle | NA | NA | 12 | 2 | 5 | 8 | 16 | 20 | 80 | 80 | 14 | 20 | 25 | NA | NA | NA | 14 | 14
Auto Flip Angle | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Mat Freq | 256 | 256 | 200 | 200 | 200 | 200 | 200 | 128 | 256 | 232 | 256 | 128 | 192 | 256 | 256 | 256 | 256 | 256
Mat Phase | 128 | 256 | 200 | 200 | 200 | 200 | 200 | 64 | 256 | 232 | 256 | 64 | 192 | 256 | 256 | 256 | 256 | 256
NEX | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | NA | 1 | 1 | 1 | 1
BW | NA | 62.5 | 31.25 | 31.25 | 31.25 | 31.25 | 31.25 | 15.63 | 31.25 | 31.25 | 27.78 | 15.63 | NA | NA | 50 | 41.67 | 27.78 | 27.78
BW2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
TI | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1297 | NA | NA | NA
INVERTF | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CV | (USERCV42=1.00) | (USERCV6=0.00)(USERCV22=0.00)(USERCV29=1.00)(USERCV30=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=1.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=1.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=1.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=1.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=1.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV2=0.00)(USERCV42=1.00) | (USERCV16=1.00)(USERCV17=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV16=1.00)(USERCV17=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=0.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV2=0.00)(USERCV42=1.00) | (USERCV0=1.00)(USERCV42=1.00) | (USERCV0=1.00)(USERCV5=1.00)(USERCV6=0.00)(USERCV9=0.00)(USERCV16=0.00)(USERCV17=1.00)(USERCV39=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV5=1.00)(USERCV30=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV22=3.00)(USERCV29=1.00)(USERCV30=1.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=0.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00) | (USERCV4=0.00)(USERCV6=0.00)(USERCV23=100.00)(USERCV40=0.00)(USERCV42=1.00)
Acceleration Phase | NA | 2 | 2 | 2 | 2 | 2 | 2 | NA | 2 | 2 | 1 | NA | 2 | NA | 2 | 2 | 1 | 1
Acceleration Slice | NA | 1 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.5 | 2 | NA | NA
Acquisition Delay | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Minimum | NA | NA | NA | NA | NA
Phases Per Location | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 40 | NA | NA | NA | NA | NA
Acquisition Order | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0 | NA | NA | NA | NA | NA
PURE | 0 | 2 | 2 | 2 | 2 | 2 | 2 | 0 | 2 | 2 | 2 | 0 | 2 | 2 | 2 | 2 | 2 | 2
Anatomical Region | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain
Coil | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head | HNS Head
CLINICALCOIL | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head | C-GE_HNS Head
COILCOMPONENT | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest | HNS Head/Neck/Chest
CoilReg | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CoilPlug | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Shim | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto
EXCITATIONMODE | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Selective | Non-Selective | Non-Selective | Selective | Selective
SL3PLANE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE1 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE2 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE3 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SLICEORDER | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SPCPERPLANE1 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE2 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE3 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PHASECORR | No | No | No | No | No | No | No | No | No | No | No | No | Yes | Yes | No | No | No | No
SAVEORIGINAL | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## FUS-Plan-BrainBBB744-HEAD
Protocol Name: | FUS-Plan-BrainBBB744-HEAD
| Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement
Field Name \ Series Name | ****CALIBRATION**** | Localizer | TGcalibration | PI-PURE calibration | ****PLANNING**** | T2w FIESTA REFORMAT | CUBE FLAIR | Axial T2-FLAIR | offlineGrid-T2STAR-SE | offlineGrid-T2STAR-ME | offlineGrid-T2STAR-ME PI | Plan-3T-BBB744-750-DV26.0-001
Series Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
Imaging Mode | 2D | 2D | 2D | 3D | 2D | 3D | 3D | 2D | 2D | 2D | 2D | 2D
Pulse Sequence | Gradient Echo | Gradient Echo | Gradient Echo | SPGR | Gradient Echo | Fiesta | Cube T2 FLAIR | FSE-XL | Gradient Echo | Multi-Echo FSPGR | Multi-Echo FSPGR | Gradient Echo
PSD Name | track | NA | B1MAP | NA | track | NA | NA | NA | NA | NA | NA | track
Imaging Options | Seq, Fast | Seq, Fast | EDR, Fast | EDR, Fast, ZIP2, Calib | Seq, Fast | NPW, EDR, Fast | EDR, Fast, ZIP512, ZIP2, T2P, IrP | FC, EDR, TRF, Fast, T2flair | FC, EDR | FC, EDR, Fast | FC, EDR, Fast, ARC | Seq, Fast
FOV | 480 | 44 | 40 | 34 | 480 | 24 | 26 | 26 | 26 | 27 | 27 | 480
Phase FOV | 1 | 1 | NA | NA | 1 | NA | 1 | 1 | 1 | 1 | 1 | 1
Slices/Slabs | 1 | 33 | 5 | 1 | 1 | 1 | 1 | 70 | 30 | 30 | 30 | 1
Scan Locs | NA | NA | NA | 32 | NA | NA | NA | NA | NA | NA | NA | NA
Slices per Slab | NA | NA | NA | NA | NA | 110 | 110 | NA | NA | NA | NA | NA
Slice Thickness | 5 | 7 | 5 | 5 | 5 | 2 | 1.6 | 2 | 3 | 3 | 3 | 5
NPW_FACTOR | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
No Slab Wrap | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slice Resolution | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Spacing | 0 | 3 | 5 | NA | 0 | NA | NA | 0 | 0 | 0 | 0 | 0
Plane | 3-PLANE | 3-PLANE | SAGITTAL | AXIAL | 3-PLANE | SAGITTAL | SAGITTAL | AXIAL | OBLIQUE | AXIAL | AXIAL | 3-PLANE
Freq DIR | Unswap | Unswap | S/I | R/L | Unswap | S/I | S/I | A/P | R/L | R/L | R/L | Unswap
TR | 20 | NA | NA | NA | 20 | NA | 4500 | 8000 | 440 | 900 | 900 | 20
Auto TR | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off
TE | Minimum | NA | NA | NA | Minimum | Min Full | Maximum | 114.4 | 20 | Min Full | Min Full | Minimum
TE2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Number of Echoes | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 6 | 6 | 1
Number of Shots | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Echo Train Length | NA | NA | NA | NA | NA | NA | 200 | 13 | NA | NA | NA | NA
Flip Angle | NA | NA | 20 | NA | NA | 55 | NA | 125 | 50 | 70 | 70 | NA
Auto Flip Angle | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Mat Freq | 1024 | 256 | 128 | NA | 1024 | 244 | 256 | 256 | 256 | 256 | 256 | 1024
Mat Phase | 2 | 128 | 128 | NA | 2 | 288 | 256 | 224 | 256 | 256 | 256 | 2
NEX | NA | 1 | 1 | NA | NA | 2 | 1 | 1 | 1 | 1 | 1 | NA
BW | 62.5 | NA | 15.63 | NA | 62.5 | 62.5 | 41.67 | 31.25 | 14.71 | 41.67 | 41.67 | 62.5
BW2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
TI | NA | NA | NA | NA | NA | NA | 1356 | 2300 | NA | NA | NA | NA
INVERTF | NA | NA | NA | NA | NA | NA | NA | Normal (A) | NA | NA | NA | NA
CV | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV2=0.00)(USERCV42=1.00) | NA | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV6=0.00)(USERCV30=1.00)(USERCV42=1.00) | (USERCV5=1.00)(USERCV8=0.00)(USERCV15=0.00)(USERCV18=1.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV38=0.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV17=1.00)(USERCV42=1.00) | (USERCV16=1.00)(USERCV17=1.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00)
Acceleration Phase | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.5 | NA
Acceleration Slice | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Delay | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Phases Per Location | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Order | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
PURE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anatomical Region | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain
Coil | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head | 3P 2ch Head
CLINICALCOIL | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head | C-FUS 2ch Head
COILCOMPONENT | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec | FUS Head and Tracking by InSightec
CoilReg | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CoilPlug | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Shim | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto
EXCITATIONMODE | Non-Selective | Selective | Selective | Selective | Non-Selective | Selective | Selective | Selective | Selective | Selective | Selective | Non-Selective
SL3PLANE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE1 | 1 | 11 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1
SL3PLANE2 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE3 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SLICEORDER | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SPCPERPLANE1 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE2 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE3 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PHASECORR | No | No | No | No | No | No | No | Yes | No | No | No | No
SAVEORIGINAL | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## FUS-Plan-BrainBBB744-BODY
Protocol Name: | FUS-Plan-BrainBBB744-BODY
| Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement | Requirement
Field Name \ Series Name | ****CALIBRATION**** | Localizer | TGcalibration | ****PLANNING**** | T2w FIESTA REFORMAT | CUBE FLAIR | Axial T2-FLAIR | offlineGrid-T2STAR-SE | offlineGrid-T2STAR-ME | Plan-3T-BBB744-750-DV26.0-001
Series Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Imaging Mode | 2D | 2D | 2D | 2D | 3D | 3D | 2D | 2D | 2D | 2D
Pulse Sequence | Gradient Echo | Gradient Echo | Gradient Echo | Gradient Echo | Fiesta | Cube T2 FLAIR | FSE-XL | Gradient Echo | Multi-Echo FSPGR | Gradient Echo
PSD Name | track | NA | B1MAP | track | NA | NA | NA | NA | NA | track
Imaging Options | Seq, Fast | Seq, Fast | EDR, Fast | Seq, Fast | NPW, EDR, Fast | EDR, Fast, ZIP512, ZIP2, T2P, IrP | FC, EDR, TRF, Fast, T2flair | FC, EDR | Fast | Seq, Fast
FOV | 480 | 44 | 40 | 480 | 24 | 26 | 26 | 25.6 | 27 | 480
Phase FOV | 1 | 1 | NA | 1 | NA | 1 | 1 | 1 | 1 | 1
Slices/Slabs | 1 | 33 | 5 | 1 | 1 | 1 | 70 | 30 | 25 | 1
Scan Locs | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slices per Slab | NA | NA | NA | NA | 110 | 110 | NA | NA | NA | NA
Slice Thickness | 5 | 7 | 5 | 5 | 2 | 1.6 | 2 | 2 | 3 | 5
NPW_FACTOR | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
No Slab Wrap | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Slice Resolution | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Spacing | 0 | 3 | 5 | 0 | NA | NA | 0 | 0 | 1.5 | 0
Plane | 3-PLANE | 3-PLANE | SAGITTAL | 3-PLANE | SAGITTAL | SAGITTAL | AXIAL | AXIAL | AXIAL | 3-PLANE
Freq DIR | Unswap | Unswap | S/I | Unswap | S/I | S/I | A/P | A/P | R/L | Unswap
TR | 20 | NA | NA | 20 | NA | 4500 | 8000 | 440 | 950 | 20
Auto TR | Off | Off | Off | Off | Off | Off | Off | Off | Off | Off
TE | Minimum | NA | NA | Minimum | Min Full | Maximum | 114.4 | 15 | Min Full | Minimum
TE2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Number of Echoes | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 8 | 1
Number of Shots | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Echo Train Length | NA | NA | NA | NA | NA | 200 | 13 | NA | NA | NA
Flip Angle | NA | NA | 20 | NA | 55 | NA | 125 | 50 | 70 | NA
Auto Flip Angle | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Mat Freq | 1024 | 256 | 128 | 1024 | 224 | 256 | 256 | 256 | 256 | 1024
Mat Phase | 2 | 128 | 128 | 2 | 256 | 256 | 224 | 256 | 256 | 2
NEX | NA | 1 | 1 | NA | 3 | 1 | 1 | 1 | 1 | NA
BW | 62.5 | NA | 15.63 | 62.5 | 62.5 | 41.67 | 31.25 | 19.23 | 41.67 | 62.5
BW2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
TI | NA | NA | NA | NA | NA | 1356 | 2300 | NA | NA | NA
INVERTF | NA | NA | NA | NA | NA | NA | Normal (A) | NA | NA | NA
CV | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV2=0.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV6=0.00)(USERCV30=1.00)(USERCV42=1.00) | (USERCV5=1.00)(USERCV8=0.00)(USERCV15=0.00)(USERCV18=1.00)(USERCV21=0.00)(USERCV22=1.00)(USERCV30=1.00)(USERCV38=0.00)(USERCV39=1.00)(USERCV42=1.00) | (USERCV42=1.00) | (USERCV16=1.00)(USERCV17=1.00)(USERCV42=1.00) | (USERCV0=2.00)(USERCV1=3.00)(USERCV2=1.00)(USERCV3=2.00)(USERCV4=1.00)(USERCV5=1.00)(USERCV6=1.00)(USERCV7=300.00)(USERCV8=300.00)(USERCV9=300.00)(USERCV10=4.00)(USERCV11=0.00)(USERCV12=0.00)(USERCV13=1.00)(USERCV14=2.00)(USERCV15=0.00)(USERCV16=0.00)(USERCV17=0.00)(USERCV42=1.00)
Acceleration Phase | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acceleration Slice | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Delay | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Phases Per Location | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Acquisition Order | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
PURE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anatomical Region | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain | Brain
Coil | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY | BODY
CLINICALCOIL | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY | C-GE_BODY
COILCOMPONENT | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil | GE T/R body coil
CoilReg | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
CoilPlug | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Shim | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto | Auto
EXCITATIONMODE | Non-Selective | Selective | Selective | Non-Selective | Selective | Selective | Selective | Selective | Selective | Non-Selective
SL3PLANE | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE1 | 1 | 11 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1
SL3PLANE2 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SL3PLANE3 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SLICEORDER | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SPCPERPLANE1 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE2 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SPCPERPLANE3 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PHASECORR | No | No | No | No | No | No | Yes | No | No | No
SAVEORIGINAL | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0